## Slide 1: On instrumental errors and related correction strategies of ozonesondes: possible effect on calculated ozone trends for the nearby sites Uccle and De Bilt
R. Van Malderen1, M.A.F. Allaart2, H. De Backer1, H.G.J. Smit3, D. De Muer1
1
2
3

## Slide 2: Testing the SI²N Ozonesonde Data Quality Assessment for the nearby stations Uccle and De Bilt
R. Van Malderen1, M.A.F. Allaart2, H. De Backer1, H.G.J. Smit3, D. De Muer1
1
2
3

## Slide 3: Outline
Motivation and introduction: O3S-DQA
The Uccle and De Bilt ozonesonde stations
Comparison of average ozone profiles
Comparison of vertical ozone trends
Conclusions

## Slide 4
Uccle & De Bilt
Motivation
Most of you are familiar with …
Average ozone profiles
Vertical ozone trends
Motivation & introduction
Conclusions
Ozonesonde observations
around 60 sites worldwide
earliest operational launches around 1965 in Resolute
“the big three”: Hohenpeissenberg (1967), Uccle (1969), Payerne (1970)
e.g. De Bilt, only since Nov. 1992

## Slide 5
The pitfalls of ozonesonde time series:
different ozonesonde types (BM, KC, ECC)
different ECC ozonesonde type manufacturers (ENSCI/DMT, Science Pump)
differences between ozonesonde production batches
different operating procedures (e.g. measurement of background current, sensing solution/buffer strengths)
different post-processing correction algorithms (e.g. normalizing the profile to the total ozone measurements at the same site)
These differences occur both between sites and at one site!
 need for homogenisation (Ozonesonde Data Quality Assessment (O3S-DQA))
Uccle & De Bilt
Average ozone profiles
Vertical ozone trends
Motivation & introduction
Conclusions

## Slide 6
standard operating procedures
guidelines for metadata collection
two standards are set: ENSCI 0.5% SST & Science Pump 1% SST (ratio is 1.0 within 1%)
transfer functions to those standards, based on double soundings/simulation chamber experiments
standard correction algorithms
uncertainty estimate for every data point: (5-6% for Uccle ECC)
(revised) worldwide, homogenous,
consistent dataset to be used for
satellite validation and trend analysis
16/04/2012
O3S-DQA principles (only for ECC):
Uccle & De Bilt
Average ozone profiles
Vertical ozone trends
Motivation & introduction
Conclusions

## Slide 7
Uccle and De Bilt
only about 175 km from each other
  horizontal ozone correlation lengths: 500 km (tropo), 1500 km (strato)
  timescales of ozone autocorrelation: 1.5-3.5 days (tropo),2-6 days (strato)
  (Liu et al. 2009, 2013)
each station represent one ECC standard
differences in operating procedures (e.g. background current measurement)
both stations apply operationally different correction strategies
unique testbed for the O3S-DQA corrections
impact of operating procedures and corrections on the (average) profile
impact of operating procedures and corrections on the trends!
Average ozone profiles
Vertical ozone trends
Conclusions
Motivation & introduction
Uccle & De Bilt

## Slide 8
very similar time variability, at different atmospheric layers
larger spread at De Bilt due to lower frequency of launches
Uccle and De Bilt
Uccle – De Bilt
Average ozone profiles
Vertical ozone trends
Conclusions
Motivation & introduction
Uccle & De Bilt

## Slide 9
differences between different Uccle correction strategies are within 3%, within 4% for De Bilt
O3S-DQA corrections: better agreement between Uccle and De Bilt only in lower stratosphere (below O3 maximum)
O3S-DQA correction at Uccle: better agreement with De Bilt above O3 maximum
Average ozone profiles
Van Malderen, R. et al., Atmos. Meas. Tech. Discuss., doi:10.5194/amt-2015-341, in review, 2016.
Motivation & introduction
Vertical ozone trends
Conclusions
Uccle & De Bilt
Average ozone profiles
1997
“relative” differences = choose a reference average profile y
(x – y)/y * 100
+ average profiles calculated from the station’s ozone profiles corrected by alternative algorithms

## Slide 10
Vertical ozone trends
differences between different Uccle correction strategies are within 3%, within 4% for De Bilt
O3S-DQA corrections: better agreement between Uccle and De Bilt only in lower stratosphere (below O3 maximum)
O3S-DQA correction at Uccle: better agreement with De Bilt above O3 maximum
trend differences between different Uccle correction strategies are within 3%/dec (UStr), within 4%/dec for De Bilt (UTr)
O3S-DQA corrections: better agreement between Uccle and De Bilt trends in LStr and LTr, but not significantly different!
O3S-DQA corrections: lower trends in Uccle, higher trends in De Bilt!
Van Malderen, R. et al., Atmos. Meas. Tech. Discuss., doi:10.5194/amt-2015-341, in review, 2016.
Motivation & introduction
Average ozone profiles
Conclusions
Vertical ozone trends
Uccle & De Bilt

## Slide 11
trend differences between different Uccle correction strategies are within 3%/dec (UStr), within 4%/dec for De Bilt (UTr)
O3S-DQA corrections: better agreement between Uccle and De Bilt trends in LStr and LTr, but not significantly different!
O3S-DQA corrections: lower trends in Uccle, higher trends in De Bilt!
Ozone recovery?
period starts with maximum peak value in EESC (1997)
only in troposphere: O3 trends significantly different from zero!
sign of the O3 trend in the strato-sphere depends on the station and on the applied data processing!
 caution is needed when using  terminology like “the onset of ozone  recovery!”
Vertical ozone trends
Van Malderen, R. et al., Atmos. Meas. Tech. Discuss., doi:10.5194/amt-2015-341, in review, 2016.
Motivation & introduction
Average ozone profiles
Conclusions
Vertical ozone trends
Uccle & De Bilt

## Slide 12
Conclusions
To serve the climate and satellite communities, an international ozonesonde data quality assessment activity was initiated.
The close ozonesonde stations Uccle and De Bilt provide a unique testbed for this homogeneity activity.
Still, natural differences in the vertical distribution of ozone between Uccle and De Bilt are likely to occur ( needs more investigation!).
The O3S-DQA corrections do not give an overall better agreement of the average profiles and trends between both stations.
The impact of the correction strategies on the trends can make the difference between the onset of ozone recovery or not.
Motivation & introduction
Average ozone profiles
Vertical ozone trends
Conclusions
Uccle & De Bilt

## Slide 13: Thank you for your attention
More details:
Van Malderen R., Allaart M.A.F., De Backer H., Smit H.G.J. and De Muer D., Atmos. Meas. Tech. Discuss., doi:10.5194/amt-2015-341, in review, 2016.

## Slide 14: Extra slides

## Slide 15
Average ozone profiles Vertical ozone trends
strong reduction of the vertical trends by the final Uccle correction w.r.t. the standard pump corrections
final Uccle correction are fairly constant and consistent: +2 to +3 % dec-1 in troposphere, -1 to -2 % dec-1 in stratosphere
large differences between the different Uccle correction steps
final Uccle correction (BM+ECC) gives differences of 8-15% compared to standard pump correction only
Van Malderen, R. et al., Atmos. Meas. Tech. Discuss., doi:10.5194/amt-2015-341, in review, 2016.

## Slide 16
Average uncertainties of the Uccle O3 profile

## Slide 17
Uccle oper. corr. (blue) reduces O3 differences to within 5%
BM sondes have a typical response equivalent to about 80% of the actual ozone amount
large differences (UT: >15%) between both De Bilt corrections
Average ozone profiles
Van Malderen, R. et al., Atmos. Meas. Tech. Discuss., doi:10.5194/amt-2015-341, in review, 2016.
Motivation & introduction
Vertical ozone trends
Conclusions
Uccle & De Bilt
Average ozone profiles

## Slide 18
O3 ↘ -0.26±0.05 %/yr (1980-1996)
O3 ↗ +0.28±0.05 %/yr (1997-2014)
full recovery at Uccle?
impact of volcanic eruptions on ozone!
other processes responsible for the O3 time variability?
Total ozone trend at Uccle

## Slide 19
Sensing solution strengths (SST)

## Slide 20
RMI conference 25/11/2015
Total ozone seasonal variation
in general, highest values in spring, lowest values in fall
Total ozone monthly trends
only in late spring-early summer (Apr-June): significant negative trends in total ozone at Uccle

## Slide 21
 stratospheric O3 ↘
 tropospheric O3 ↗
BUT: different periods = different ozonesonde types!
= BM & ECC
Vertical ozone distribution at Uccle

## Slide 22
 stratospheric O3 ↘
at around -2%/decade
 tropospheric O3 ↗
at around +2%/decade
Impact of the different ozonesonde types on the trends?
Vertical ozone distribution at Uccle